# MAINTAINCE SHEET FOR GC-NET v3

| Station name: | NSE |  | Date: | 01-05-2026 |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| People present: | RSF, RABNI, IASTE, ARU |  | Maintenance |  | Install | Take-down | Data Download |

| Weather: | Sunny, high clouds |
| --- | --- |

| METADATA |  |  |
| --- | --- | --- |
|  | Arrival | Departure |
| Download Data | YES / NO | NA |
| Time difference between logger and UTC (± mm:ss) | : | : |
| Screenshot of ‘Status’ table | YES / NO | YES / NO |
| Screenshot of ‘Public’ table in FastScan mode | YES / NO | YES / NO |
| Photos of mast, loggerbox and sensors (inc. radiometer serial) | YES / NO | YES / NO |

| STATION POSITON AND OREINTATION |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
|  |  |  | Arrival |  | Departure |
| Lower lufft height (see A) |  |  | 188 cm |  | 300 cm |
| Lower SR50 height (see A) |  |  | 145 cm |  | 265 cm |
| Upper lufft height (see A) |  |  | 320 cm |  | 425 cm |
| Upper SR50 height (see A) |  |  | 260 cm |  | 380 cm |
| Upper boom direction (towards radiometer, see B) |  |  | 338 degree mag. |  | 338 degree mag. |
| Lower boom direction (towards lufft, see B) |  |  | 248 degree mag. |  | 248 degree mag. |
| Upper wind sensor orientation (away from black box, see C) |  |  | Yes |  | Yes |
| Lower wind sensor orientation (away from black box, see C) |  |  | Broken |  | Yes |
| Radiometer across-boom tilt, looking at the front (see D) |  |  | 4 degree to right/east |  | NA |
| Mast tilt north-south axis |  |  | 5 degree down-north |  | 0 degree |
| Mast tilt east-west axis |  |  | 1 degree to right/east |  | 2 degree right/east |
| Height of thermistor string pipe above surface |  |  | 28 cm |  | 230 cm |
| Depth of snowboard | NA | Height of snowboard stake |  | NA Vertical: YES / NO |  |

| Station name: | NSE |  | Date: | 01-05-2026 |
| --- | --- | --- | --- | --- |

| STATION HARDWARE |  |  |  |
| --- | --- | --- | --- |
| To do? | Done? | Item | Comments |
| Update program | yes | Loggerbox | ID arrival: M3 – 03 ID departure: M3-03 |
|  |  | Modem IMEI | IMEI arrival: 300534063818800 IMEI departure: 300534063818800 |
|  |  | Radiometer | Serial arrival: 244328 Serial departure: 244328 |
| Replace | Yes | Wind sensor upper | Type arrival: Black / White Type departure: Black / White |
| Replace | Yes | Wind sensor lower | Type arrival: Black / White Type departure: Black / White |
| Replace | Yes | Lufft upper |  |
| Replace | Yes | Lufft lower |  |
| Replace | Yes | SR50 upper |  |
| Replace | Yes | SR50 lower |  |
|  |  | Thermistor string |  |
|  |  | Iridium antenna |  |
|  |  | Solar panel |  |
| Replace | Yes | Battery box | Burial under lower lufft: YES / NO Location if NO: |
| Replace | Yes | Mast |  |

NOTES (e.g. sensor damage, must-do at next visit)

We did not find old snowboard stake New snowboard top 250 cm No liquid GNSS installed New loggerbox for next year: Unstable USB connection port on logger New solar panel: Plug is damaged.