# MAINTAINCE SHEET FOR GC-NET v3

| Station name: | SDL |  | Date: | 07-05-2026 |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| People present: | RSF, HETJA, IASTE, ARU, ASNOR |  | Maintenance |  | Install | Take-down | Data Download |

| Weather: | Sunny, high clouds |
| --- | --- |

| METADATA |  |  |
| --- | --- | --- |
|  | Arrival | Departure |
| Download Data | YES / NO | NA |
| Time difference between logger and UTC (± mm:ss) | 0:12 | 0:0 |
| Screenshot of ‘Status’ table | YES / NO | YES / NO |
| Screenshot of ‘Public’ table in FastScan mode | YES / NO | YES / NO |
| Photos of mast, loggerbox and sensors (inc. radiometer serial) | YES / NO | YES / NO |

| STATION POSITON AND OREINTATION |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
|  |  |  | Arrival |  | Departure |
| Lower lufft height (see A) |  |  | 275 cm |  | 255 cm |
| Lower SR50 height (see A) |  |  | 246 cm |  | 246 cm |
| Upper lufft height (see A) |  |  | 371 cm |  | 371 cm |
| Upper SR50 height (see A) |  |  | 400 cm |  | 400 cm |
| Upper boom direction (towards radiometer, see B) |  |  | 324 degree mag. |  | 324 degree mag. |
| Lower boom direction (towards lufft, see B) |  |  | 234 degree mag. |  | 234 degree mag. |
| Upper wind sensor orientation (away from black box, see C) |  |  | Yes |  | Yes |
| Lower wind sensor orientation (away from black box, see C) |  |  | Yes |  | Yes |
| Radiometer across-boom tilt, looking at the front (see D) |  |  | 2.5 degree to west |  | 0 degree |
| Mast tilt north-south axis |  |  | 4 degree down-north |  | 4 degree down-north |
| Mast tilt east-west axis |  |  | 2 degree to left/west |  | 2 degree to left/west |
| Height of thermistor string pipe above surface |  |  | 115 cm |  | 115 cm |
| Depth of snowboard | 110 cm | Height of snowboard stake |  | 58 cm Vertical: YES / NO |  |

| Station name: | SDL |  | Date: | 07-05-2026 |
| --- | --- | --- | --- | --- |

| STATION HARDWARE |  |  |  |
| --- | --- | --- | --- |
| To do? | Done? | Item | Comments |
| Update program | yes | Loggerbox | ID arrival: M3 – 06 ID departure: M3-06 |
|  |  | Modem IMEI | IMEI arrival: 300534063815680 IMEI departure: 300534063815680 |
| Replace | Yes | Radiometer | Serial arrival: 213250 Serial departure: 110560 |
| Replace | Yes | Wind sensor upper | Type arrival: Black / White Type departure: Black / White |
| Replace | Yes | Wind sensor lower | Type arrival: Black / White Type departure: Black / White |
| Replace | Yes | Lufft upper |  |
| Replace | Yes | Lufft lower |  |
| Replace | No | SR50 upper | Reused old sensor |
| Replace | Yes | SR50 lower |  |
| Replace | Yes | Thermistor string | Installed new thermistor string |
|  |  | Iridium antenna |  |
|  |  | Solar panel |  |
|  |  | Battery box | Burial under lower lufft: YES / NO Location if NO: |
| Replace | no | Mast | No need |

NOTES (e.g. sensor damage, must-do at next visit)

We had no extension cable for thermistor, so we installed a new in the old tube. Tube was not extended because cable was too short. Bamboo stick and carbonfiber stake marks the top of the thermistor tube. New snowboard top 195 cm Liquidice GNSS installed