AWS maintenance / establishment checklist

| Station name | MIT |
| --- | --- |
| People present | MCIT,MAV |
| Purpose of visit | Maintenance of AWS |
| Date & time | Arrival: 01.09.2013 08:11 Departure: 01.09.13 |
| Transport | By: Helicopter, Greenland Air From: Tasiilaq |
| Weather | Cloudy / Sunny |

Remarks on state of AWS / maintenance / establishment

Metadata upon arrival (Establishment of new AWS: Go to last page)

| Time difference between logger clock and UTC | sec |
| --- | --- |
| Adjust logger time to UTC (be sure PC runs on UTC time) | Y/N |
| Download data to PC and/or change CF Card (wait for green light) | Both Y/N |
| Name of logger program | 2012a |
| Photo or screen dump of values in fast scan mode | Y/N |
| Latitude (dd mm.mmm) |  |
| Longitude (dd mm.mmm) |  |
| Altitude (m) |  |
| Photos of tripod, sensors, logger box wiring etc. (more is better) | Y/N |
| Mast tilt in boom direction (+ if radiometer tilting down) | o |
| Mast tilt across boom looking from radiometer (- if clockwise) | o |
| Boom direction relative to north | ? |
| Wind sensor box direction along boom? If not, measure direction. | Y/N |
| Radiometer aligned with mast and inclinometer? If not, measure. | Y/N |
| Temperature / humidity height (bottom of casing – surface) | cm |
| Sonic ranger on AWS height (membrane – surface) | cm |
| Wind sensor height (center of propeller – surface) | cm |
| Sonic ranger on stakes height (membrane – surface) | cm |
| Free length of stake – outer, holding boom with sonic ranger | cm |
| Free length of stake – middle | cm |
| Free length of stake – outer, holding boom without sonic ranger | cm |
| Free length of ablation hose on/above the ice surface | Not fully visible m |
| Alt: depth of ablation hose using markings on hose | Min. 8 m N/A |
| Height of ablation hose fluid level above the ice surface | cm |
| Vertical difference in surface height (at station – at hose) (approx) | cm |
| Length of thermistor string on surface from surface marking | cm |
| Height of surface irregularities | 50 cm |
| Stake assembly still standing? | Y/N |
| Apparent damage to sensors? If yes, which? | Y/N |

Remarks:

Tripod severely damaged probably from being buried in snow during the winter spring/season. Boom and legs heavily bended and the Letrona insert stucked to the boom. No apparent damage to the sensors. 
Serial numbers and maintenance details

| Part | Old part number | New part number, if replaced | Potential maintenance tasks |
| --- | --- | --- | --- |
| Radiometer | 080165 | 131307 | - Aligned with mast: N - Clean: Y |
| Inclinometer | 00658 | 00415 | - Aligned with radiometer within 0.5: N |
| Satellite antenna | No SN DB: GEUS08 |  | Not shifted |
| Wind sensor | 79485 | 124021 |  |
| Temperature / humidity assembly (Fan, Bracket, MP100, Hygroclip) | ASS 2 / 2009 |  | - New HygroClip with old assembly: - Fan spins and sounds OK: Y Note Ass no. + serial no. for all 4 parts ! |
| Or: HygroClip only | S3: 60091546 MP100H: 60276145 DB: RS12T 60123277 DB: MKRS 60276140 | H2-S3: 61012684 MP102H: 61154940 RS12T: 61148916 MKRS: 61154943 | - Large offset in HygroClip temperature: Y/N |
| Sonic ranger on AWS | 1147 (?) DB: 1141 | “ | - Old sensor with new membrane: Y/N - Clicking as it should: Y |
| Solar panel | 5436470 | “ | - Output OK: Y - Clean: Y |
| GPS antenna | DB: 018111213 |  | - Stuck to top of logger enclosure: N |
| Data logger | WP E3202 22108 |  | - New internal battery in old logger: N - New operating system: N - New logger program: 2012a.cr1 |
| Card reader | 4590 |  | - LED either green of flashing orange (active): Y |
| Multiplexer | E4787 |  | - Clicking as it should: Y |
| Iridium modem IMEI number | 2259 830 DC-DC: GEUS16 |  | - Transmissions tested: N - Problems: Y/N |
| Barometer | 3766881 | “ |  |
| Sonic ranger on stakes | 3501 | 5110 | - Old sensor with new membrane: Y(?) - Clicking as it should: Y |
| Thermistor string | DB: B09 | D B16 | - New string in use, but old one still in ice: Y |
| Ablation hose assembly | 5019ABS | “ | - New hose in use, but old one still in ice: Y/N - Air in current hose: Y/N - Air removed from current hose: Y/N - Added 50% antifreeze mix to bladder: Y/N |
| Logger enclosure | - Replaced (including everything inside): N - New vent filter in logger enclosure: N - Moisture inside: N - New desiccant bags (2): Y |  |  |
| Battery box | - New box with new batteries: N - Old box with new batteries: N - Moisture inside: N - New desiccant bags (2):N, - not opened |  |  |
| Stakes | - New ones drilled in: Y - Length: 8 m |  |  |
| Tripod | - New parts: Y all new tripod + wires except Letrona. The latter alu-insert damaged but working. |  |  |

- Metadata after maintenance / establishment

| Time difference between logger clock and UTC | 0 min |
| --- | --- |
| Adjust logger time to UTC (be sure PC runs on UTC time) | Y |
| Download data to PC and/or change CF Card (wait for green light) | Y |
| Name of logger program | Promice2012b.CR1 |
| Photo or screen dump of values in fast scan mode | Y |
| Latitude (dd mm.mmm) |  |
| Longitude (dd mm.mmm) |  |
| Altitude (m) |  |
| Photos of tripod, sensors, logger box wiring etc. (more is better) | Y |
| Mast tilt in boom direction (+ if radiometer tilting down) | 2 |
| Mast tilt across boom looking from radiometer (- if clockwise) | 0 |
| Boom direction relative to north (radiometer should point to true south) | 30 magn. |
| Wind sensor box direction along boom? If not, change. | Y |
| Radiometer aligned with mast and inclinometer? If not, align. | Y |
| Temperature / humidity height (bottom of casing – surface) | 290 cm |
| Sonic ranger on AWS height (membrane – surface) | 290 cm |
| Wind sensor height (center of propeller – surface) | 335 cm |
| Sonic ranger on stakes height (membrane – surface) | 35 cm |
| Free length of stake – outer, holding boom with sonic ranger | 65 (41*) cm |
| Free length of stake – middle | 65 (50) cm |
| Free length of stake – outer, holding boom without sonic ranger | 68 (50) cm |
| Free length of ablation hose on/above the ice surface | NOT VISIBLE |
| Alt: depth of ablation hose using markings on hose | NOT VISIBLE |
| Height of ablation hose fluid level above the ice surface | 150 cm |
| Vertical difference in surface height (at station – at hose) (approx) | - cm |
| Length of thermistor string on surface from surface marking | -65 cm |
| Height of surface irregularities | 20 cm |

Remarks:  Red mark on thermistor string went to depth of 65cm. * free length in () marks height from surface to horizontal stake.