AWS maintenance / establishment checklist

| Station name | TAS_U |
| --- | --- |
| People present | MCIT,MAV, Marek |
| Purpose of visit | Maintenance of AWS |
| Date & time | Arrival: 29.08.2013 14:54 Departure: 29.08.13 |
| Transport | By: Helicopter, AS 350 Greenland Air From: Tasiilaq |
| Weather | Overcast |

Remarks on state of AWS / maintenance / establishment

Metadata upon arrival (Establishment of new AWS: Go to last page)

| Time difference between logger clock and UTC | sec |
| --- | --- |
| Adjust logger time to UTC (be sure PC runs on UTC time) | Y/N |
| Download data to PC and/or change CF Card (wait for green light) | Both Y/N |
| Name of logger program |  |
| Photo or screen dump of values in fast scan mode | Y/N |
| Latitude (dd mm.mmm) |  |
| Longitude (dd mm.mmm) |  |
| Altitude (m) |  |
| Photos of tripod, sensors, logger box wiring etc. (more is better) | Y/N |
| Mast tilt in boom direction (+ if radiometer tilting down) | -4o |
| Mast tilt across boom looking from radiometer (- if clockwise) | +1o |
| Boom direction relative to north |  |
| Wind sensor box direction along boom? If not, measure direction. | Y/N |
| Radiometer aligned with mast and inclinometer? If not, measure. | Y/N |
| Temperature / humidity height (bottom of casing – surface) | 257 cm |
| Sonic ranger on AWS height (membrane – surface) | 270 cm |
| Wind sensor height (center of propeller – surface) | 315 cm |
| Sonic ranger on stakes height (membrane – surface) | 295 cm |
| Free length of stake – outer, holding boom with sonic ranger | 290 cm |
| Free length of stake – middle | 300 cm |
| Free length of stake – outer, holding boom without sonic ranger | 304 cm |
| Free length of ablation hose on/above the ice surface | m |
| Alt: depth of ablation hose using markings on hose | N/A |
| Height of ablation hose fluid level above the ice surface | cm |
| Vertical difference in surface height (at station – at hose) (approx) | cm |
| Length of thermistor string on surface from surface marking | cm |
| Height of surface irregularities | 50 cm |
| Stake assembly still standing? | Y/N |
| Apparent damage to sensors? If yes, which? | Y/N |

Remarks:

Station standing between crevasses but in general in Good shape and without difficulties to access the station. Stake assembly slightly out of order and only boom without SR not connected to stakes.

Serial numbers and maintenance details

| Part | Old part number | New part number, if replaced | Potential maintenance tasks |
| --- | --- | --- | --- |
| Radiometer | 080164 | “ | - Aligned with mast: Y - Clean: Y |
| Inclinometer | 000088 | 000089 | - Aligned with radiometer within 0.5: Y |
| Satellite antenna | GEUS12 | “ |  |
| Wind sensor | 110446 | “ |  |
| Temperature / humidity assembly (Fan, Bracket, MP100, Hygroclip) | ASS 6/2011 RS12T: 60123285 S3: 55024226 MP100: 60179402 MKRS: 60276143 (fra Ass6/2011) |  | - New HygroClip with old assembly: Y - Fan spins and sounds OK: Y Note Ass no. + serial no. for all 4 parts ! |
| Or: HygroClip only |  | S3: 51382185 | - Large offset in HygroClip temperature: Y/N |
| Sonic ranger on AWS | 1736 |  | - Old sensor with new membrane: Y/N - Clicking as it should: Y |
| Solar panel | ? |  | - Output OK: Y - Clean: Y |
| GPS antenna |  | ? (036110634 DB) | - Stuck to top of logger enclosure: Y |
| Data logger | WP E3201 22107 | “ | - New internal battery in old logger: N - New operating system: Y - New logger program: 2013c.cr1 |
| Card reader | 4442 | “ | - LED either green of flashing orange (active): Y |
| Multiplexer | DB: E4796 | “ | - Clicking as it should: Y |
| Iridium modem IMEI number | - | 1218 580 | - Transmissions tested: Y - Problems: N |
| Barometer | 3814092 | “ |  |
| Sonic ranger on stakes | 3873 | “ | - Old sensor with new membrane: Y - Clicking as it should: Y |
| Thermistor string | R31 (DB) |  | - New string in use, but old one still in ice: Y/N |
| Ablation hose assembly | 596ABS (DB) |  | - New hose in use, but old one still in ice: N - Air in current hose: N - Air removed from current hose: N - Added 50% antifreeze mix to bladder: N |
| Logger enclosure | - Replaced (including everything inside): N - New vent filter in logger enclosure: N - Moisture inside: N - New desiccant bags (2): Y |  |  |
| Battery box | - New box with new batteries: N - Old box with new batteries: N - Moisture inside: N - New desiccant bags (2):N, - not opened |  |  |
| Stakes | - New ones drilled in: Y - Length: 8 m |  |  |
| Tripod | - New parts: N |  |  |

- Metadata after maintenance / establishment

| Time difference between logger clock and UTC | 0 min |
| --- | --- |
| Adjust logger time to UTC (be sure PC runs on UTC time) | Y |
| Download data to PC and/or change CF Card (wait for green light) | Y |
| Name of logger program | Promice2013c.CR1 |
| Photo or screen dump of values in fast scan mode | Y |
| Latitude (dd mm.mmm) |  |
| Longitude (dd mm.mmm) |  |
| Altitude (m) |  |
| Photos of tripod, sensors, logger box wiring etc. (more is better) | Y |
| Mast tilt in boom direction (+ if radiometer tilting down) |  |
| Mast tilt across boom looking from radiometer (- if clockwise) |  |
| Boom direction relative to north (radiometer should point to true south) | 32 magn. |
| Wind sensor box direction along boom? If not, change. | Y |
| Radiometer aligned with mast and inclinometer? If not, align. | Y |
| Temperature / humidity height (bottom of casing – surface) | cm |
| Sonic ranger on AWS height (membrane – surface) | cm |
| Wind sensor height (center of propeller – surface) | cm |
| Sonic ranger on stakes height (membrane – surface) | 60 cm |
| Free length of stake – outer, holding boom with sonic ranger | 76 cm |
| Free length of stake – middle | 93 cm |
| Free length of stake – outer, holding boom without sonic ranger | 88 cm |
| Free length of ablation hose on/above the ice surface | PART NOT VISIBLE m |
| Alt: depth of ablation hose using markings on hose | PART NOT VISIBLE cm |
| Height of ablation hose fluid level above the ice surface | 169 cm |
| Vertical difference in surface height (at station – at hose) (approx) | +50 cm |
| Length of thermistor string on surface from surface marking | Top 5 marks + |
| Height of surface irregularities | 50 cm |

Remarks: